Laser System Standard Operating Procedure (SOP)

# SCST 287

This Laser Standard Operating Procedure has been written to explain the safe use of the laser(s) identified below.  Laser Supervisors and Laser Users must follow this procedure. When the Laser (see section 1 below) is in use, everyone working in the lab must wear approved PPE, which consists solely of laser safety goggles that have an optical density of 5 and work in the range of wavelengths of the laser beam. These laser goggles are stowed on the interior side of the entry way door. Before the laser can be used, the laser curtain that passively shields all those who inadvertently look into the lab through the window in the entry door from the hallway, as well as those who walk into the lab and have not yet put on the goggles, must be in place. Note: IF THE LASER IS ON, a notice to that effect is posted on the lab entry door that opens to the hallway beforehand.

# LASER SAFETY CONTACTS for the Laboratory for Basic Plasma Physics

| Supervisor: | Greg Severn | Phone: | X6845 | Mobile: | 8586033087 |
| --- | --- | --- | --- | --- | --- |
| Primary Laser User: | Greg Severn | Phone: | X6845 | Mobile: | 8586033087 |
| Laser Safety Officer: | Keith MacDonald | Phone:: | 619-260-4211 | Mobile: | 6194582594 |
| Medical Emergencies and Fire: | Public Safety | Phone: | 2222 |

- LASER DESCRIPTION

| Laser Type: | Tunable Diode Laser | Laser Location |
| --- | --- | --- |
| Manufacturer: | New Focus (Newport) | Department: | Physics & Biophysics |
| Model: | Vortex model 6000 | Building: | SCST |
| Class: | x 3B 4 | Room Number: | 287 |

| Emission | Wavelength(s) or Wavelength Range (nm) | Power (mW) | Pulse Energy (J) | Pulse Duration (sec) | Pulse Frequency (Hz) | Eyewear Optical Density Required |
| --- | --- | --- | --- | --- | --- | --- |
| x CW Pulsed | 729-729 | 30 | n/a | n/a | n/a | 5 |

# Brief Description of Laser Use (specific to the lab):

- SETUP, ALIGNMENT AND OPERATING PROCEDURES

- Alignment/Setup
(Please indicate who will be responsible for performing these procedures and include specific beam alignment/visualization aids to be used as well as PPE.)

Dr. Severn is solely responsible for set-up and alignment for all tunable diode lasers in the Plasma Lab.

- Start-up and Operation:
      Before turning on the system for the first time, check that the AC line voltage indicator on the back of the controller matches the voltage you are using (see page 55 of Vortex manual in the manuals subdirectory of D:/PaX_Research).

To turn on the Stable Wave/Vortex:

1.Turn the power key switch on the front panel clockwise (to the “|” position).

This turns on AC power for the entire laser system, including the temperature circuit in the laser head.

2.Wait at least 20 minutes after turning on the key switch to allow the laser head to reach a stable temperature.

3.Press the Power button on the controller front panel to allow current to flow to the laser head. The button will flash during the five -second safety delay before the current is activated. The button will remain lit while current is flowing to the laser diode.

4. Choose “Constant-Current” mode by toggling the “mode button” on the front panel of the Vortex Laser Controller.

The diode will not emit laser light until the current is set high enough to meet the threshold current.

Before turning off the system, you should first turn off power to the laser diode by pressing the Power button located on the front panel.

- Shutdown
To turn off the Stable  Wave/Vortex:

- Press the Power button on the controller front panel to shut off current flow to the laser head

- Turn off the key switch to power down the Laser and controller.

| Laser Hazard Control Measures |
| --- |
| Check if applicable | Control | Comments |
|  | Entryway (door) interlocks |  |
|  | Laser protective housing interlocks |  |
|  | Emergency stop/panic button |  |
|  | Master switch (operated by key or computer password) |  |
| x | Beam stops/attenuators | The only beam stops are walls and curtains. Within the Laser Hazard zone |
|  | Beam path enclosure (e.g., light pipe) |  |
| x | Protective barriers | Laser Hazard zone created by curtains that must remain in place, surrounding the optics table. Laser safety goggles are stored on the inside of the entry door, as shown in Fig. 1 below. As the curtain DOES NOT partition the room, but only affords safe entry into it (see map below). All students and staff that need to go beyond the curtain during their visit to the lab, MUST wear approved Laser goggles WHEN THE LASER IS ON. |
| x | Warning signs |  |
| x | Warning lights |  |
|  | Other (specify): |  |

PERSONAL PROTECTIVE EQUIPMENT

# Laser Protective Eyewear

| wear this eyewear |
| --- |
| Eyewear Manufacturer | Protected Wavelength(s) | Optical Density (OD) |
| ThorLabs (LG4) | 630-850nm | >5 |

- Non-Beam Hazards of this System (check all that apply)

Check all non-beam hazards that apply and provide a brief description of the control measure(s) implemented to control the hazard.

Chemical (dyes, solvents, etc.); attach applicable Safety Data Sheet (SDS)

Biological agents

Electrical (high voltage, large current, capacitors, etc.)

Laser Generated Air Contaminants

Compressed gases or cryogenic liquids

Fire/ignition source

Plasma/blue light exposure

Other (specify):

# Emergency procedures

(Post in the lab)

# In Case of Emergency

- Shut down laser (if it is safe to do so)

- If there is a fire or medical emergency, call the Public Safety 2222.

- For non-emergency injuries, call Public Safety at 7777 and they will escort students to the Student Health Center or check with the staff member to take them to the Workman’s Comp. clinic.

- Call the Plasma Lab Laser Supervisor and the Laser Safety Officer:

- LASER SAFETY CONTACTS for the Laboratory for Basic Plasma Physics

| Supervisor: | Greg Severn | Phone: | X6845 | Mobile: | 8586033087 |
| --- | --- | --- | --- | --- | --- |
| Primary Laser User: | Greg Severn | Phone: | X6845 | Mobile: | 8586033087 |
| Laser Safety Officer: | Keith MacDonald | Phone:: | 619-260-4211 | Mobile: |  |
| Medical Emergencies and Fire: | Public Safety | Phone: | 2222 |

# Laser User Review:

# I have read and understand this procedure and have been trained on implementing its contents.

| Name (Printed) | Signature | Date |
| --- | --- | --- |

# Text boxes

Fig. 1 Schematic of lab and photo of view from the entry door of the lab (SCST 287). Stay behind laser curtains for passive safety. However. If the Laser is ON, a note is to be posted on the door, then please put on a pair of laser safety goggles. Find one in the pouches fastened to the inside of the entry way door itself.